## Reset IC
MCU NRST Voltage detection threshold
| VDD | 3.3
| VIH | 1.8665
| VIL | 1.06
Reset IC operation
| Vcc | /RESET
| 3.3 | 3.3
| 3.2 | 3.2
| 3.1 | 3.1
| 3 | 3
| 2.9 | 2.9
| 2.8 | 2.8
| 2.7 | 2.7
| 2.6 | 0
NRST의 High level이 1.87이므로 /RESET 2.7V도 High로 인식한다.
23ㅛㅓ

## RS-232
LF Gen V1.0 |  |  | LF Gen Origin
MCU |  | B-to-B | B-to-B | Max232
USART1_RX | PA10 | 24 | 24 | R1_OUT | 12
USART1_TX | PA9 | 22 | 22 | T1_IN | 11
Pipette Proto V1.0 |  |  |  | Pipette Proto V2.0
MCU |  | Max3232 |  | MCU |  | Max3232
USART1_RX | PA9 | R1_OUT | 12 | USART1_RX | PA10 | R1_OUT | 12
USART1_TX | PA10 | T1_IN | 11 | USART1_TX | PA9 | T1_IN | 11

## SW Review
No | Issue Date | Status | Closed Date | Issue | Command
1 | 2018-06-27 |  |  | HAL_GPIO_WritePin(GPIOB, PWR_LED_Pin¦PLA_ON_LED_Pin, GPIO_PIN_RESET);
2
3
4
5
6
7
8
9
10
11
gpio.c
main.c
Clitask/device_rtx_cmds.c

## Power
LF Generator Power Set-up
Measure
|  | Frequency[KHz] | Power [W] | Pulse width [usec] | On-Time [usec]
|  | 50 | 10 | 2 | 82
|  | 50 | 20 | 2 | 182
|  | 50 | 100 | 2 | 900
|  | Frequency[KHz] | Power [W] | Pulse width [usec] | On-Time [usec]
|  | 50 | 20 | 2 | 182
|  | 30 | 20 | 3.2 | 70
Target Power에 따른 On Time 계산
Pulse의 Frequeny가 가변되더라도 Pulse Duty가 동일하기 때문에 On Time에 영향이 없다.
|  |  | Simulation | V1.0 | V1.0 | V1.0 | Origin | Unit
|  | Vin | 220 | 220 | 220 | 220 | 220 | VAC
|  | Vout | 300 | 300 | 311.127 | 300 | 220 | VDC | 220VAC는 RMS value이며, 정류시 Peak Voltage이 311VDC가 된다.
|  | Vtran_in | 150 | 150 | 155.5635 | 150 | 220 | VDC | Half Bridge 방식으므 Vout/2기준으로 Vout과 0V로 스위칭 된다.
|  | Itran_in | 18.8 | 18.8 | 20 | 12.5 | 12.5 | A
|  | P_peak | 2820 | 2820 | 3111.2698 | 1875 | 2750 | W
|  | Turn ratio | 5 | 5 | 5 | 5 | 5 | N2/N1
|  | Vtran_out | 720 | 750 | 777.8175 | 750 | 1100 | VDC
|  | Itran_out | 3.67 | 3.76 | 4 | 2.5 | 2.5 | A
|  | Pulse Duty | 10 | 10 | 10 | 10 | 10 | %
|  | Total Duty | 20 | 20 | 20 | 20 | 20 | % | Pulse가 1주기당 2ea가 발생하므로 Total Duty는 Pulse Duty x 2 이다.
|  | P_real | 564 | 564 | 622.254 | 375 | 550 | W | 실제 전체 Power = P_peak*(2xPulse_Duty)/100
|  | P_target | 300 | 300 | 300 | 300 | 300 | W
|  | Power ratio | 0.53191 | 0.53191 | 0.48212 | 0.8 | 0.54545 | P/P | P_target / P_real
|  | On Time | 2.6596 | 2.6596 | 2.4106 | 4 | 2.7273 | msec | P_target / P_real x 5msec
|  | Pulse Frequency | 50 | 50 | 50 | 50 | 50 | KHz
|  | Pulse Period | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | msec
|  | Pulse No | 132.9787 | 132.9787 | 120.5296 | 200 | 136.3636
=> Origin은 220VAC rms value로 계산했기 때문에 맞지 않음
=> 실측 Itran_out에 맞춰 Power를 계산하여 V1.0과 비슷한 결과가 나온것임.
Limit [msec]
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5
5

## Power 설정
On Time calculator | [5 empty] | Duty 조정
| P_peak[W] | 2750 | 2750 | [5 empty] | TIM Clock
| Pulse Duty | 10 | 10 |  |  | Frequency | 80 | KHz | SystemCoreClock | 48000000 | HCLK
| P_real[W] | 550 | 550 |  |  | T | 12.5 | usec | PCLK1 | 48000000 | HCLK / APB1 prescaler
| FULL SPAN | 5 | 100 | msec |  | Duty | 10 | % | TIM clock | 48000000 | PCLK1 x 1
|  | 200 | 10 | Hz |  |  |  |  | TIM counter clock | 16000000
| P_target[W] | On Time[msec] | On Time[msec] | [5 empty] | Prescaler | 2 | (TIM clock /TIM counter clock) - 1
| 20 | 0.18182 | 3.6364 |  | =($C11/E$7)*E$8 |  |  |  | TIM output clock | 80000
| 40 | 0.36364 | 7.2727 |  | (Setup Power / Max Power * SPAN) |  | 24.8 | usec | period (ARR) | 199 | (TIM counter clock / TIM output clock) - 1
| 60 | 0.54545 | 10.9091 | [5 empty] | Duty | 50
| 80 | 0.72727 | 14.5455 | [5 empty] | Pulse | 99.5
| 100 | 0.90909 | 18.1818 | [5 empty] | TIM output Clock T | 1.25e-05
| 120 | 1.0909 | 21.8182
| 140 | 1.2727 | 25.4545
| 160 | 1.4545 | 29.0909
| 180 | 1.6364 | 32.7273
| 200 | 1.8182 | 36.3636
| 220 | 2 | 40
| 240 | 2.1818 | 43.6364
| 260 | 2.3636 | 47.2727
| 280 | 2.5455 | 50.9091
| 300 | 2.7273 | 54.5455
| 320 | 2.9091 | 58.1818
| 340 | 3.0909 | 61.8182
| 360 | 3.2727 | 65.4545
| 380 | 3.4545 | 69.0909
| 400 | 3.6364 | 72.7273
| 420 | 3.8182 | 76.3636
| 440 | 4 | 80
| 460 | 4.1818 | 83.6364
| 480 | 4.3636 | 87.2727
| 500 | 4.5455 | 90.9091
| 520 | 4.7273 | 94.5455
| 540 | 4.9091 | 98.1818
| 560 | 5.0909 | 101.8182
| 580 | 5.2727 | 105.4545
| 600 | 5.4545 | 109.0909
| 620 | 5.6364 | 112.7273
| 640 | 5.8182 | 116.3636
| 660 | 6 | 120
| 680 | 6.1818 | 123.6364
| 700 | 6.3636 | 127.2727
| 720 | 6.5455 | 130.9091
| 740 | 6.7273 | 134.5455
| 760 | 6.9091 | 138.1818
| 780 | 7.0909 | 141.8182
| 800 | 7.2727 | 145.4545
| 820 | 7.4545 | 149.0909
| 840 | 7.6364 | 152.7273
| 860 | 7.8182 | 156.3636
| 880 | 8 | 160
| 900 | 8.1818 | 163.6364